## Slide 1: Quantitative Optical Coherence Elastography as an Inverse Elasticity Problem
Presenter: Li Dong (PhD student)
Rensselaer Polytechnic Institute
February 14, 2016, Optics Elastography and Tissue Mechanics III, BiOS, Photonics West, SPIE.
Authors: Li Dong (RPI), Philip Wijesinghe (UWA), James Dantuono (RPI), David Sampson (UWA), Peter Munro (UCL, UWA), Brendan Kennedy (UWA), and Assad Oberai (RPI)
Notes: Spell out names, phd student

## Slide 2: Outline
Elastography background
Inverse elasticity problem
Solution strategy
Experimental validation cases for quantitative OCE
Summary

## Slide 3: Elastography background
Estimate disp., such as image cross correlation
Imaging modalities, such as ultrasound, OCT, MRI, etc.
Differentiate
Strain
Normalized Young’s modulus
Inverse Problem
US B-mode (Hall, UW)
Liu, Oberai, et al. PloS one, 2015.

## Slide 4: Inverse elasticity problem
Tikhonov
TV
Tikhonov
Total Variation (TV)
Add noise
u
ũ
Apply regularization

Direct method
Optimization method

?
?
Quasi-static equilibrium for linear elasticity:
Notes: Explain all variables

## Slide 5: Solution strategy
PDE-constrained optimization
Two PDE constraints
BCs for the PDE constraints
Find the shear modulus distribution μ that minimizes the cost function:
subject to
where

## Slide 6: Solution strategy cont.
Quasi-static
Linear elasticity
Incompressible
Plane strain in 2D
Calibration layer
Sample layer
Ux=0, Uy=ũ (and Uz=0)
Traction free on x (and z), Uy=ũ
BCs:

## Slide 7: Solution strategy cont.: adjoint method
subject to
Naïve way
Adjoint way
where
MxM * MxN
MxN
MxM * Mx1
Mx1

## Slide 8: Case I: Tissue-like phantom in 2D
Displacement
Strain
Shear modulus
OCT
9.3 kPa vs 10.6 kPa
1.9 kPa vs 2.2 kPa
500μm
52,150 optimization variables, terminate after 2,525 iterations.
Notes: 300x350=105000 points, 982x3500 microns, axial by lateral.
For the inclusion case, it takes about 17 hours to finish the case with optimal regularization parameter(3e-3), terminating at 2525th iteration.
All results are obtained from the linux platform with 8-core 2.3GHz Opteron processor and 16 GB memory.
Erase arrows

## Slide 9: Case II: Equine bronchus in 2D
M: mucosa, ASM: airway smooth muscle, C: cartilage
Displacement
Strain
Shear modulus
OCT
500μm
1.9 kPa
26,379 optimization variables, terminate after 6,000 iterations.
Dong, L., Wijesinghe, P., Dantuono, J.T., Sampson, D.D., Munro, P.R.T., Kennedy, B.F., and Oberai, A. A. “Quantitative Compression Optical Coherence Elastography as an Inverse Elasticity Problem.” IEEE Journal of Selected Topics in Quantum Electronics. 2016. In press.
Dong, L., et al.. “Quantitative Compression Optical Coherence Elastography as an Inverse Elasticity Problem.” IEEE Journal of Selected Topics in Quantum Electronics. 2016. In press.
Notes: 160x440=70400 points, 524x4400 microns, axial x lateral.
For the tissue case, it takes about 30 hours to finish the case with optimal regularization parameter(1e-3), terminating at max (6000) iteration.
All results are obtained from the linux platform with 8-core 2.3GHz Opteron processor and 16 GB memory.

## Slide 10: Case III: Tissue-like phantom in 3D
In the OCT experiment,
31.5 million points are sampled in a regular grid.
This results in 157.5 million tetrahedron elements and 15.75 million optimization variables.
Too much and rather unnecessary to work with!
μm
Displacement field ũ
Notes: Explain displacement plots

## Slide 11: Case III: Tissue-like phantom in 3D cont.
Remeshing ( )
Interpolation
Continuation
# | Points | Elements | Optimization variables
Regular mesh | 31,500,000 | 157,500,000 | 15,750,000
Adaptive mesh | 60,091 | 365,525 | 54,607
Adaptive Mesh
Unpublished work.
Notes: Remesh the elements that have grad(Mu) > threshold and are below the calibration layer
Interpolate nodal values between different meshes
Use the solution of the previous step as the initial guess of the next step
We deterine mesh size in the next to be inverse proportional to the gradi of mu

## Slide 12: Case III: Tissue-like phantom in 3D cont.
1.9 kPa vs 2.3 kPa
1.9 kPa
9.3 kPa vs 10.1 kPa
54,607 optimization variables, terminate after 470 iterations.
Unpublished work.
Notes: 300x350x300 points, 982x3500x3000 microns, axial by lateral by elevational.
For the 3D inclusion case, it takes about 35 hours with regularization parameter(3e-3) and terminates at 470th iteration.
All results are obtained from the linux platform with 8-core 2.3GHz Opteron processor and 16 GB memory.

## Slide 13: Summary
Demonstrate a quantitative OCE method
  Compression OCE
  Calibration layer
  Optimization method
  Proper regularization
  Adjoint equations
  Adaptive mesh
Validate the method against a tissue-like phantom with stiff inclusion in 2D and 3D
Apply the method to an equine bronchus sample
Future work: investigate nonlinear effects, assumptions on boundary conditions, in-vivo tests, etc.

## Slide 14: Acknowledgement
Funding agencies: NSF and the National Breast Cancer Foundation in US; University of Western Australia, Australian Research Council and the National Health and Medical Research Council in Australia; Royal Society in UK.
Authors: Li Dong (RPI), Philip Wijesinghe (UWA), James Dantuono (RPI), David Sampson (UWA), Peter Munro (UCL, UWA), Brendan Kennedy (UWA), and Assad Oberai (RPI)
Questions?